# ROTRONIC HygroClip Digital Input / Output

OEM customers that use the HygroClip have the choice of using either the analog humidity and temperature output signals or the digital signal input / output (DIO).

Using the digital signal output offers the following benefits:

( Higher resolution: this is of special interest when working with temperature in (F. Internally, the HygroClip reads the signals from both the humidity and temperature sensor with 16-bit resolution. A resolution of 0.004%rh and 0.004(C is available with the digital output of the HygroClip. By contrast, the D/A conversion used to generate the analog outputs is limited to a maximum resolution of 11 to 12 bits. (0.025 to 0.05%rh and 0.03 to 0.06(C).

( Simpler processing of the signals provided by the HygroClip regarding humidity computations, data recording, etc.

( No requirement to do an initial calibration of the circuits used to read the HygroClip

( Maximum flexibility regarding both the measuring range and the engineering units of the output system.

( Generally more reliable communication between the HygroClip and other devices

# 1. General Description of the HygroClip

The ROTRONIC HygroClip is a humidity-temperature probe which plugs into any matching connector from ROTRONIC. The HygroClip has 5 pins (or wires) corresponding to the following:

( Supply Voltage (+) : 3.5 to 50 VDC

( Ground (-) : reference for supply voltage and output signals

( Humidity Analog Output (+)

( Temperature Analog Output (+)

( DIO: digital input – output

## 1.1 Internal Subsystems of the HygroClip

Internally. the HygroClip is comprised of two main subsystems:

( AIRCHIP 2000

This ASIC includes the circuitry required to measure the capacitive humidity sensor and the Pt100 RTD and to convert the measurements into digital counts. The ASIC also includes two D/A converters that convert the data from the microcontroller into analog output signals. The ASIC also regulates all supply voltages and generates the reset and clock for the microcontroller.

( Microcontroller / EEPROM

The microcontroller uses the counts measured by the ASIC to compute the value of humidity and temperature. Calibration data, linearization and other sensor data are memorized in the EEPROM. The microcontroller sends data both to the DIO pin (digital output signal) and to the ASIC (analog output signal).

Two way communication between the HygroClip and an external system (for example, a PC) is used to write data to the EEPROM such as calibration data, serial number, etc.

## 1.2 Operating Modes of the HygroClip

The HygroClip does not require an initialization sequence or “send” command. About 3 seconds after being powered up, the HygroClip automatically sends the first humidity and temperature data and updates all outputs every 0.66 second (measuring cycle).

Essentially, the HygroClip has two operating modes: the send mode (or normal operating mode) and the receive mode.

In the send mode, the HygroClip operates as a conventional humidity and temperature probe with two analog outputs and one digital output. Unless other instructions have been received (see recive mode), the data from both the Pt100 RTD temperature sensor and the capacitive humidity sensor are sent during each measuring cycle (0.66 sec) to both the digital and analog outputs.

In the receive mode, digital commands can be sent to the HygroClip to the purpose of calibrating the HygroClip or to ask for additional information such as the serial number, calibration date, etc.

## 1.3 The HygroClip One-Wire Digital Interface

Because of the limited number of connections available on the HygroClip, all digital communication between the HygroClip and an external system is done with a single DIO pin that is referenced to ground.

The main characteristics of the HygroClip communications protocol are as follows:

( Bit Definition:

The communication protocol differenciates between logical “1” and logical “0” bits based on the amount of time (pulse width) between two successive transitions. Communication is done without grouping the data into bytes. As opposed to this, communication is a single bit stream with a very short pause between each bit.

( 2-way Communication

Communication is entirely defined by means of time intervals or windows. The HygroClip always assumes the function of master and signals the start of each communication cycle (generally every 0.66 sec.). At the beginning of each communication cycle, the HygroClip goes into the “receive” mode and the external system has a defined time window during which it can begin sending a bit stream. If no bit stream is detected within this time window, the HygroClip automatically switches back to the “send “mode and proceeds with sending the humidity and temperature data.

( Transmission Speed

The pulse and pause times were designed to permit a communication frequency of 1.6 MHz. Data is exchanged at a rate of about 4 to 5 ms per byte.

( Integrity of the Output Data

The protocol makes use of a check sum only for the output of the humidity and temperature data. For all other data, only the number of bits is verified (must be a multiple of 8) and overall verification of data integrity should be done by the external system.

# 2. Specifications of the Digital Input Output (DIO)

## 2.1 General

The digital output of the HygroClip (DIO) allows the bi-directional transmission of data (input / output). The electrical connection consists of a single wire plus ground.

Data is transmitted by means of a stream of pulses with a nominal height of 3.15 V, referred to ground. The current to and from the DIO should always be with the range of (300(A.

Any circuit used to read the DIO ((C) should have a high input impedance. The circuit should also include a pull down resistor to the purpose of defining the voltage of the DIO during the short intervals of time when it is floating (receive mode). The pull down resistor is not required when the receiving circuit uses a transistor inverter to convert the pulse height from 3.15 V to 5.0 V.

## 2.2 Definition of a Data Bit

The HygroClip uses a single DIO pin (referred to ground) for sending and receiving digital data. Pulse width is used to generate or read a “1” or a “0” data bit.

When inactive, the state of the DIO is a logical 1. The transmission of each individual bit (logical 0 or 1) requires from 400 (s to 540 (s (470 (s nominal). The transmission of each bit begins with a negative transition of the DIO state from a logical 1 to a logical 0. This transition (A) is used as the time origin.

Based on the above diagram, nominal times are as follows:

| tAB | minimum width of a “1” bit | 80(s |
| --- | --- | --- |
| tAC | maximum width of a “1” bit | 115(s |
| tBC | time window during which a transition is permitted | 35(s |
| tAD | minimum width of a “0” bit | 240(s |
| tAE | maximum width of a “0” bit | 325(s |
| tDE | time window during which a transition is permitted | 85(s |
| tAF | Next “reference” transition – earliest time | 400(s |
| tAF | Next “reference” transition – latest time | 540(s |
| tEF | minimum duration of the pause | 145(s |

## 2.3 Description of the Communication Cycle

The start of each communication cycle is set by the HygroClip. The HygroClip microcontroller begins by generating a “0” bit. The negative transition of this bit (S) is used as the time origin. The “0” bit has a minimum duration of 400(s and a maximum duration of 540(s. After this, the HygroClip marks a pause (I - R) and lets the DIO float (receive mode) during the time window (R - O). This allows an external device to set the logical state of the DIO.  Before transmitting the first bit, the external device must set the DIO to a logical “1”. In the receive mode, the HygroClip monitors the logical state of the DIO and looks for a negative transition within the time window (R – O).

( Case 1: no negative transition is detected within the (R – O) window:

The HygroClip takes control of the DIO and sets it to a logical “1” after 800(s. After a maximum of 5,500(s, the HygroClip starts sending the humidity and temperature data. At this time, the HygroClip generates a negative transition (A) indicating the beginning of the first bit of the temperature and humidity data stream. Each byte is sent with the least significant bit (LSB) first and the most significant bit (MSB) last.

( Case 2: a negative transition is detected within the (R – O) window:

As mentioned earlier, the external device sets the DIO to a logical “1”. Transmission of the first bit by the external device begins with a negative transition (after 640(s at the earliest and 800(s at the latest). After detecting the negative transition, the HygroClip stays in the “receive” mode and takes the incoming bit stream, LSB first and MSB last.

If at any time, the HygroClip does not detect a negative transition over a time period of 540(s (maximum duration of a “0” bit), it automatically returns to the send mode.

During a 2-way communication, the HygroClip does not automatically send the temperature and humidity data. The data sent by the HygroClip depends on the commands received from the external device. Powering the HygroClip off and on restores normal operation (send mode).

## 2.4 Tolerance to Distortions in the Transmission Line

The different times mentioned earlier are measured directly at the output of the sender. Logical levels of the DIO are defined as <1/3 and >2/3 of the nominal voltage.

Protective circuits in the transmission line as well as signal amplifiers and level adapters can also cause slight changes in the width of the pulses during the transmission of data. The ROTRONIC communication protocol allows the following maximum tolerances regarding the width of a pulse (time elapsed between the negative and positive transitions): -30(s and +15(s.

For the HygroClip as well as for any external device communicating with the HygroClip, this means that the following times should be acceptable:

HygroClip:

| tAB | minimum width of a “1” bit | 50(s |
| --- | --- | --- |
| tAC | maximum width of a “1” bit | 130(s |
| tBC | time window during which a transition is permitted | 80(s |
| tAD | minimum width of a “0” bit | 210(s |
| tAE | maximum width of a “0” bit | 340(s |
| tDE | time window during which a transition is permitted | 130(s |
| tAF | Next “reference” transition – earliest time | 370(s |
| tAF | Next “reference” transition – latest time | 555(s |
| tEF | minimum duration of the pause | 100(s |

External Device:

| tSI | minimum width of the “0” bit | 370(s |
| --- | --- | --- |
| tSI | maximum width of the “0” bit | 555(s |
| tSR | earliest start of the bit stream | 610(s |
| tSO | latest start of the bit stream | 820(s |

# 3. Encoding of the Temperature and Humidity Data

## 3.1 Structure of the Data String

In the normal mode, the HygroClip sends every 0.66 second a temperature and humidity data string of constant length to the DIO. The data string is made of 7 bytes
(8 bits per byte). For each byte, the least significant bit (LSB) is sent first and the most significant bit (MSB) is sent last.

The first 3 bytes are used for temperature, the next 3 bytes for humidity and the last byte is a checksum.

The temperature range of the HygroClip is –50..200(C. In order to eliminate the need for transmitting the minus sign in the case of negative temperatures, the digital value of temperature is offset by +50(C and has a range of 0..250(C. The 50(C offset should be subtracted when computing temperature from its digital value.

The following table shows the structure of the data string (the symbol ( indicates any hexadecimal value):

| Byte # | Hex Value | Description |
| --- | --- | --- |
| 1 | 0x54 | ASCII character T |
| 2 | 0x(( | Decimal portion of temperature [(C / 256] **) |
| 3 | 0x(( | Non decimal portion of temperature [(C] |
| 4 | 0x46 | ASCII character F |
| 5 | 0x(( | Decimal portion of humidity [%rh / 256] **) |
| 6 | 0x(( | Non decimal portion of Humidity [%rh] |
| 7 | 0x(( | Checksum |

The checksum is defined as the sum of the bytes 1 to 6, modulo 256 (the remainder of the division of the value of the sum by the number 256).

**) for an explanation, see example below

## 3.2	Example of a Data String

In this example, the HygroClip is sending the following binary data string (starting from the left).

00101010110001010100010001100010001000000011101011111101

As already mentioned, the least significant bit (LSB) of each byte is sent first and the most significant bit (MSB) is sent last. Using the LSB first rule, the binary data sting can be converted as follows, :

Hexadecimal

| T | (C dec | (C whole | F | %rh dec | %rh whole | ChkSum |
| --- | --- | --- | --- | --- | --- | --- |
| 0x54 | 0xA3 | 0x22 | 0x46 | 0x04 | 0x5C | 0xBF |

Decimal

| 84 | 163 | 34 | 70 | 4 | 92 | 191 |
| --- | --- | --- | --- | --- | --- | --- |

The decimal value 84 corresponds to the ASCII character T and the decimal value 70 corresponds to the ASCII character F.

The checksum can be verified as follows:

(84 + 163 + 34 + 70 + 4 + 92) = 447 mod 256 = 191 (OK)

Temperature can be computed as follows:

(163 / 256) + 34 – 50 = -15.637 (C

Humidity can be computed as follows:

(4 / 256) + 92 = 92.016 %rh

Note that the use of 1 byte (8 bits or 256 counts) for the transmission of decimal values provides a resolution of 1 / 256 or 0.004(C or 0.004 %rh.

# 4. Special Applications

The HygroClip can be calibrated or adjusted, and additional information can be obtained, by sending the appropriate series of commands.

Very few OEM applications should require the use of the 2-way communication capability of the HygroClip.

## 4.1 Items Accessible by 2-way Communication

- Raw A to D counts for the temperature channel (resulting from the resistance value of the Pt100 RTD, the value of the reference resistor and the offset voltage).

- Raw A to D counts for the humidity channel (a pair of values that are averaged)

- Read and adjust the data used for converting resistance into temperature (used for calibration).

- Read and adjust the data used for converting capacitance into humidity (used for calibration)

- Freeze the analog outputs to a constant value and return the outputs to their normal mode (normal operation is also automatically restored when the HygroClip is powered off and on).

- Set the digital output value of the temperature channel

- Set the digital output value of the humidity channel

- Set or adjust the conversion of digital values into analog signals for the temperature channel

- Set or adjust the conversion of digital values into analog signals for the humidity channel

- Read the serial number of the HygroClip, the software version number, the date of the latest calibration. Read the data associated with the nominal characteristic curve of the sensors or the compensation.

## 4.2 Technical Support from Rotronic

On request, OEM customers can get from ROTRONIC an example of a simple circuit that converts the digital signal from the HygroClip into the RS232 format for direct use with the serial port of a PC.

Customers who wish to make use of the 2-way communication feature of the HygroClip should contact ROTRONIC or their local representative for additional information and support.

C51程序供参考

/****the softwire for oem010711 ****/

#include <reg51.h>

#include <math.h>

#define WREN 6

#define WRDI 4

#define RDSR 5

#define WRSR 1

#define READ0 3

#define READ1 11

#define WRITE0 2

#define WRITE1 10

#define EEPSR 0

#define DEVICE_NUMBER 0x00         //device number

#define CRC_CONSTANT 0xa001

/////////////for com

#define VER "000000"

#define OEM1 "    JWSH-WD     "

#define OEM2 "     V1.0       "

#define TEM_ZERO_H 0xf0

#define TEM_ZERO_L 0x60

#define TEM_FULL_H 0x17

#define TEM_FULL_L 0x70

#define TEM_HIGH_H 0x11

#define TEM_HIGH_L 0x94

#define TEM_LOW_H 0xfe

#define TEM_LOW_L 0x0c

/////////for display

#define HW_DIS1 "    T= 00.0C    "

#define HW_DIS2 "    H= 00.0%    "

#define ADDR_SETUP   "1:Input Comm_addr"

#define ADDR_NUM     "  Comm_addr=000  "

#define SAVE_CH_ADDR "  Saving.......  "

#define NO_CHANGE    " Not save change "

/*********  for system  ************/

sbit lcdled=P0^7;

sbit db7=P0^6;

sbit db6=P0^5;

sbit db5=P0^4;

sbit db4=P0^3;

sbit lcde=P0^2;

sbit lcdrw=P0^1;

sbit lcdrs=P0^0;

//sbit alarm=P1^0;

//sbit ledrun=P1^1;

//sbit dincs1=P1^2;

sbit a4051a=P1^3;

sbit a4051b=P1^4;

sbit a4051c=P1^5;

sbit ledrun=P1^6;  /* if sipcs1=0 then x25045 run else analog 4051 run*/

sbit kcom=P1^7;

sbit dec=P2^7;

sbit menu=P2^6;

sbit ent=P2^5;

sbit add=P2^4;

sbit out4=P2^3;

sbit out2=P2^2;  /*if outcs=0 then output 74hc245 relay data */

sbit out3=P2^1;

sbit out1=P2^0;

sbit rx=P3^0;

sbit tx=P3^1;

sbit dd=P3^2;

sbit sipin=P3^3;

sbit sipclk=P3^4;

sbit sipcs1=P3^5;

sbit sipout=P3^6;

sbit alarm=P3^7;

/* for communication */

static unsigned char r_buffer[15];

static unsigned char password,com_addr,com_boud,r_pointer,clear_word;

bit break_flag,com_flag,tel_send_flag,com_error_flag,alm_change_flag,relay_flag,com_setup_flag;

/*main*/

int tem_value[4],hum_value[4],temvalue,humvalue,tem_high,

tem_low,hum_high,hum_low,tem_correct,hum_correct,tem_full,

tem_zero,hum_full,hum_zero;

unsigned char dd_number[7],relay;

bit alm_tem_h,alm_hum_h,alm_tem_l,alm_hum_l;

unsigned int com_num,int1=0;

static unsigned char lcd_com,lcd_data;

static char *dis_lcd;

unsigned int crc_result;

static unsigned char eep_addr,eep_data;

static bit eep_flag;      /* if eep_flag=0, write or read low part */

main()

{

void dog();

void delay();

void timer0();

char i,n;

int int1;

bit test_flag,out_flag,return_flag=0;

P0=0xff;

P1=0xff;

P2=0xff;

P3=0xff;

IE=0;

IP=0x00;   //规定串口中断为高优先级中断源

SCON=0x70;

TMOD=0x21;

TH1=0xfd; //

TL1=0xfd;

PCON=0x10; // pcon.4 is power flag

TR1=1;

ET1=0;

ES=1;

break_flag=0;

test_flag=0;

eep_flag=0;

eep_addr=0;

out_flag=0;

n=0;

com_setup_flag=0;

for (;;)

{

//     read_eep();

for (;1;)

{

if (dd==0) break;

dog();

}

timer0();

delay();

if (!break_flag) {for(i=0;i<7;i++) dd_number[i]=0x00;}

else

{

tem_value[0]=tem_value[1];

tem_value[1]=tem_value[2];

tem_value[2]=tem_value[3];

tem_value[3]=((dd_number[1]*100/256)+dd_number[2]*100-5000);

hum_value[0]=hum_value[1];

hum_value[1]=hum_value[2];

hum_value[2]=hum_value[3];

hum_value[3]=((dd_number[4]*100/256)+dd_number[5]*100);

n++;

if (n>8) {out_flag=1;n=10;}

else {out_flag=0;}

}

if (out_flag)

{

temvalue=((tem_value[0]+tem_value[1]+tem_value[2]+tem_value[3])/4)+tem_correct;

humvalue=((hum_value[0]+hum_value[1]+hum_value[2]+hum_value[3])/4)+hum_correct;

dis_input();//dis_input();

if (com_setup_flag)

{

write_eep();

com_setup_flag=0;

}

an_th();

out_send();

alarm=~alarm;

if (!EA) EA=1;

}

}

}

void dog()

{

sipcs1=0;

sipcs1=1;

}

void delay()

{

int delay_number;

for (delay_number=0;delay_number<30000;delay_number++) dog();

}

void timer0()  // interrupt 0

{

unsigned int number;

char i,k;

unsigned char m;

bit start_flag=0,error_flag;

//  TMOD=0X01;

for (;!start_flag;)     // the start of data

{

TR0=1;

for (;!dd;)

{

dog();

}

TR0=0;

if (TH0>3) start_flag=1;

else start_flag=0;

}

for (;dd;) dog();

error_flag=0;    // recieve data

for (k=0;k<7;k++)

{

for (i=0;i<8;i++)

{

TH0=0;TL0=0;

TR0=1;       /* open timer0*/

for (;!dd;)

{

dog();

}

TR0=0;		/* close timer0*/

number=(TH0*256+TL0);

m>>=1;

if ((number<150)&&(number>50))  m+=128;

else if ((number>200)&&(number<400)) error_flag=0;

else error_flag=1;

if (error_flag==1) break;

for (;dd;) dog();

}

if (!error_flag) dd_number[k]=m;

}

if (!error_flag)    // check data

{

m=0;

for (i=0;i<6;i++) m+=dd_number[i];

if (m==dd_number[6]) {error_flag=0;break_flag=1;}

else error_flag=1;

}

dog();

}

汇编语言(ASM文件)程序供参考

ORG	0000H

LJMP	AA

ORG	000BH

LJMP	BSX

ORG	0013H

LJMP	DINT1	    ;External-1

ORG	001BH

LJMP	DT1	    ;Timer-1

********

；定时器0中断

BSX:	MOV	TH0,#0F8H   ;2ms定时

MOV	TL0,#0D7H

JNB	26H,DDS1     ；BIT26H是正在处理接受到的数据

CLR	EX1          ；如果正在处理，此时禁止响应外部中断和定时器1中断

CLR	ET1

LJMP	DDSE

DDS1:	SETB	EX1          ；否则打开外部中断和定时器1中断

SETB	ET1

DDSE:	RETI

；外部中断

DINT1:	JNB	26H,DINT2   ；BIT 26H是否已经开始接收

CLR	ET1       ；关闭定时器1

RETI

DINT2:	SETB	ET1       ；允许定时器1中断并赋初值170uS定时值，开始计时

MOV	TL1,#65H

MOV	TH1,#0FFH

SETB	TR1

MOV	56H,#00H

RETI

；定时器1中断

DT1:	MOV	TL1,#65H   ；

MOV	TH1,#0FFH

MOV	2EH,A      ；将ACC，PSW，R0，放入2EH，2FH，30H，中进行保护

MOV	A,PSW

MOV	2FH,A

MOV	30H,R0

JB	27H,DT2    ；BIT，27H是数据开始接收标志，=1则到DT2

INC	56H        ；否则累计到否170uS*21=3.5mS(初始标志为5mS)

MOV	A,56H

CJNE	A,#15H,DT11

DT12:	CLR	TR1        ;如果已经接收到有效初始标志位,则停止定时器1中断,等待数据位的外部中断1有效再打开定时器1

CLR	ET1

SETB	27H         ;置数据接收开始标志

MOV	57H,#08H    ;应接收7字节数据

MOV	58H,#00H    ;从第一字节第零位开始

LJMP	DT1E

DT11:	JNC	DT12

LJMP	DT1E

DT2:	CLR	TR1          ;已经开始接收数据,此时中断说明当前位数据已经有效,等待下一位数据开始产生外部中断再打开定时器1中断

CLR	ET1          ;

MOV	56H,#00H

MOV	A,58H

ADD	A,#59H       ;接收到数据放在59H,5AH,5BH,5CH,5DH,5EH,5FH

MOV	R0,A

MOV	A,@R0        ;找到当前接收的数据位置

MOV	C,P3.3       ;判断此位是1或0放入数据保存位置

RRC	A

MOV	@R0,A

DEC	57H           ;当前正在接收字节已接收位数减1

MOV	A,57H

JNZ	DT1E

MOV	57H,#08H      ;当前字节接收完成则准备接收下一字节

INC	58H

MOV	A,58H

CJNE	A,#07H,DT1E   ;如果已经接收到7个字节则置数据处理标志和清开始接收标志

DT4:	CLR	27H

SETB	26H

DT1E:	MOV	R0,30H    ;恢复被保护内容

MOV	A,2FH

MOV	PSW,A

MOV	A,2EH

RETI

;主程序

AA:	CLR	A

MOV	56H,A

MOV	57H,A

MOV	58H,A

MOV	59H,A

MOV	5AH,A

MOV	5BH,A

MOV	5CH,A

MOV	5DH,A

MOV	5EH,A

MOV	5FH,A

A2:	MOV	TMOD,#11H

MOV	TH0,#0F8H

MOV	TL0,#0CDH

MOV	IP,#0CH     ;置外部中断1和定时器1中断为高优先级

MOV	IE,#92H

A3:	SETB	TR0         ;打开定时器1和外部中断

CLR	TR1

SETB	IT1

SETB	EX1

JNB	26H,DDAE     ;是否可以开始处理接收到数据

MOV	58H,#00H

MOV	57H,#08H

MOV	R2,#06H

MOV	R0,#59H

MOV	R1,#00H

DDA1:	MOV	A,@R0        ;接收到数据进行累加计算校验和

INC	R0

ADD	A,R1

MOV	R1,A

DJNZ	R2,DDA1

MOV	A,R1

CJNE	A,5FH,DDAE

CLR	C            ;校验和正确

MOV	A,5AH	 ;5A/25=X.0

MOV	B,#19H

DIV	AB

MOV	5AH,A

MOV	A,5BH

MOV	B,#0AH    ;计算温度值

MUL	AB

ADD	A,5AH

MOV	5BH,A

MOV	A,B

ADDC	A,#00H

MOV	5AH,A

MOV	R1,5AH

MOV	R2,5BH	   ;T-50.0

MOV	R3,#01H

MOV	R4,#0F4H

LCALL	SUBS       ;R1,R2-R3,R4子程序,结果=R1,R2

MOV	R0,#0DCH

MOV	A,R1

MOV	@R0,A

INC	R0

MOV	A,R2

MOV	@R0,A

MOV	A,5DH      ;计算湿度

MOV	B,#19H

DIV	AB

MOV	5DH,A

MOV	A,5EH

MOV	B,#0AH

MUL	AB

ADD	A,5DH

MOV	5EH,A

MOV	A,B

ADDC	A,#00H

MOV	5DH,A

MOV	R0,#0DEH

MOV	A,5DH

MOV	@R0,A

INC	R0

MOV	A,5EH

MOV	@R0,A

DDAE:	CLR	26H

# Text boxes

DIO

DIO

“0” bit

“1” bit

A

B

C

D

E

F

0

80(s

115(s

240(s

325(s

470(s

min 400(s

max 540(s

DIO

5500(s

maximum

0

470(s

min 400(s

max 540(s

640(s

800(s

S

I

R

O

A

Receive Mode (DIO floats)

Start of first output bit

(if no negative transition during R-O)

R-O: time window available to the external device to initiate the communication